## Sheet1
Version | Desc | Test_Acc | Test_Loss | tcSeizures | allSeizures | FalseAlarms | Notes
OSD_v1 | Original OSD Algorithm | - | - | 0.78 | 0.73 | 0.61 | This is the current algorithm used by OpenSeizureDetector with default settings.
V0.01 | First Go – all seizure data | 0.86 | 0.35 | - | - | -
V0.02 | -20 to +40s of event used | 0.92 | 0.3 | 0.76 | 0.74 | 0.82 | OSD V4.1.0a &b: Good detection performance, but false alarmed when walking.
V0.03 | Additional data | 0.95 | 0.14 | 0.43 | 0.23 | 1 | Poor seizure detection performance so not used
V0.04 | Random Oversample | 0.903 | 0.277 | 1 | 0.98 | 0.9 | OSD V4.1.0c: Good false alarm performance, but difficult to simulate a seizure to make it alarm - is it trained to be too specific to the seizures in the training set?
V0.05 | SMOTE Oversample | 0.87 | 0.29 | 0.97 | 0.98 | 0.86 | OSD V4.1.0d: Good false alarm and detection performance against testdataset.
V0.06 | More normal data | 0.88 | 0.28 | 0.95 | 0.96 | 0.87 | OSD V4.1.0e:
V0.07 | More normal data | 0.87 | 0.3 | 0.97 | 0.98 | 0.82 | OSD V4.1.0f
V0.08 | More normal data | 0.87 | 0.32 | 0.89 | 0.89 | 0.92 | not used because spoted some bad seizure data in database
V0.09 | Removed bad seizure data | 0.92 | 0.19 | 1 | 1 | 0.94 | OSD V4.1.0g